## Slide 1: IQ to B-mode Conversion & Endocardium Segmentation in Parasternal Short-axis Echocardiograms
Addison Elliott
04/18/2018

## Slide 2: What are Echocardiograms?
Ultrasound imaging of the heart (image called B-mode)
Notes: So you have a transducer that you place on the subject and it sends sound waves out in a fanning direction.
It’s taking images across time as well.
So a B-mode image is just one still image from the transducer and a B-mode loop is the entire time frame.
Show B-mode video.
This B-mode loop is taken near the apex or bottom of the heart.
Show where endocardium is.

## Slide 3: What are Parasternal Short-axis views (PSAX)?
View near apex of heart, just a donut
Ignore Plane A

## Slide 4: Project Purpose
Part 1
  Convert raw transducer data into B-mode images that you saw
  Allows tuning of parameters (DR, TGC, gain, etc)
  No public source code, have to start from scratch
Part 2
  Use B-mode images at short-axis to segment short-axis views

## Slide 5: Part 1: IQ to B-mode Conversion

## Slide 6: How Ultrasound Works
Transducer emits sound waves at frequencies of 1-10MHz (ultrasound)
Transducer measures amplitude and delay of reflected waves
Raw signal from transducer is radiofrequency (RF) data
  In-phase and quadrature phase (IQ) data is similar

## Slide 7: Block Diagram

## Slide 8: Review

## Slide 9: Speckle Reduction
Very difficult problem to solve, many existing methods out there
In this study, following methods were examined:
  Gaussian filter
  Perona-Malik anisotropic diffusion
  Median filter
  Lee filter
  Wavelet denoising

## Slide 10: No Noise Reduction
Notes: Smoothing filter applied to image

## Slide 11: Gaussian Filter
Notes: Smoothing filter applied to image

## Slide 12: Anisotropic Diffusion
Notes: Smoothing filter applied to image

## Slide 13: Median Filter
Notes: Smoothing filter applied to image

## Slide 14: Lee Filter
Notes: Smoothing filter applied to image

## Slide 15: Wavelet’s Denoising
Notes: Smoothing filter applied to image

## Slide 16: Ultrasound Machine
Notes: Smoothing filter applied to image

## Slide 17: Key Takeaways
Denoising is a second order problem
  Small changes
  Last 10% of work takes 90% of time
Gaussian & Anisotropic diffusion blur image
  Not great for cardiologists, great for computer algorithms
Median, Lee & Wavelet
  Visually about the same
  Chose Wavelet because it seems slightly better than others

## Slide 18: Part 2: Segmenting Endocardium

## Slide 19: B-Mode PSAX near apex
Red- Center of left ventricle (LV)
Blue – Endocardial boundary

## Slide 20: B-Mode PSAX at mitral valve
Red- Center of left ventricle (LV)
Blue – Endocardial boundary
Green – Mitral valve

## Slide 21: B-Mode PSAX at papillary muscles
Red- Center of left ventricle (LV)
Blue – Endocardial boundary
Green – Papillary muscles

## Slide 22: Why do we care?
According to AHA, coronary heart disease is the #1 cause of death globally.
Measures to help diagnose and preemptively treat
  Left Ventricle Ejection Fraction (LVEF)
  Left Ventricle Mass/Volume
Ultrasound is primary method of cardiac health assessment
  Cheap, portable, and non-invasive
  MRI is bulky, time-consuming scans and expensive
  Downside is ultrasound is very noise, called speckle noise
Tedious to do this manually, not typically done in practice.
Create segmentation algorithm to automatically calculate measures
Eventually quantify fat around heart

## Slide 23: Segmentation Method
Smooth image to remove speckle noise (Anisotropic diffusion)
Transform to polar domain with center user-specified
Get edge image to use as cost
  Sobel filter along radial direction
  Equalize histogram
Perform dynamic programming using Dijkstra’s algorithm
  Start point = user-specified start point
  End point = one pixel to the left of start point
  For each point, can only move East and adjacent pixels
  Wrap around in the horizontal direction (angles are periodic)
  Cost function = Edge image (inverted)

## Slide 24: Smoothed Image

## Slide 25: Polar Conversion
Notes: Explain that X-axis is angular dimension, Y-axis is radial dimension. Origin is at lower-left

## Slide 26: Endocardial Boundary (Apex)

## Slide 27: Endocardial Boundary (Mitral)
Notes: Missing boundary, weak boundaries. Messing up

## Slide 28: Additional Cost Functions
Missing or weak boundaries (like Mitral valve), does not work
Absolute Difference Method
  Take radial difference from start point to current point, add to cost function
  Constrains to circle with radius of start point
Derivative Method
  Take derivative from current point to 5th optimal point back, add to cost function
  Smooths out path, makes more circular

## Slide 29: Mitral Valve w/ Abs Diff

## Slide 30: Mitral Valve w/ Derivative

## Slide 31: Key Points
Did not get a chance to run on PSAX papillary muscle view
Messes up with missing or weak boundaries
Absolute Difference Cost Function
  Did not work well with filling in missing boundaries
Derivative Cost Function
  Did not work well with filling in missing boundaries
  Worked nicely to smooth the contour however
Neither cost function could fill in missing boundaries

## Slide 32: Future Work
Factor in temporal dimension
  Use gradient for cost function
Use shape model or ASMs to segment boundaries
Consider Kalman filter to leverage multiple methods

## Slide 33: Conclusion
Created algorithm to convert raw data to B-mode image
Created segmentation algorithm to segment endocardial boundary
  Works well enough if boundaries are clearly defined, not always case for ultrasound
Build on both as future work, combine them to have one process.
  Allows most parameters to tune

## Slide 34: Questions?

## Slide 35: Literature Review
Zhang et al. employed dynamic programming (DP) using Dijkstra’s algorithm on cardiac MRI image
Numerous shape-based models in echocardiograms
  Good for interpolating missing boundaries
  Requires existing shape model to be made
Active shape models are another common method
  Start with initial contour and slowly move towards gradient of cost function
  Factor in curvature and continuity

## Slide 36: Project Purpose
B-mode image is missing a lot of information
E.g. spectral parameters of RF data can be analyzed in regions to differentiate tissues (Akhila’s project)
Need to have the B-mode image and RF data
  Select region from B-mode
  Get corresponding RF data from that region
Cannot collect both B-mode and RF data from machine.
This is why I am converting the IQ data to B-mode images
Also, allows tuning of parameters (DR, TGC, gain)

## Slide 37: Related Works
Most ultrasound machines have proprietary software to convert from RF to B-mode image
Few papers on conversion from RF to B-mode image, but no speckle reduction
Few papers focusing on speckle reduction only
Importantly, no papers with publicly accessible source code. No code for us to use!

## Slide 38: Software
Python 3.5.4 using Anaconda
Libraries:
  NumPy
  SciPy
  Scikit-image
  matplotlib
  PyQt5
  Jupyter for debugging

## Slide 39: B-mode images
Raw data is converted into a brightness mode (B-mode) image
This is what most people see when dealing with ultrasound
Short-axis view of heart
Ultrasound is noisy (low SNR) Noise is termed ‘speckle’

## Slide 40: Envelope Detection
IQ data is complex
To make complex data real, envelope detection is performed.
This is also just taking the magnitude of the complex data.

## Slide 41: Log Compression
Reduces dynamic range of the data so that it is better displayed on a monitor
Has the effect of scaling down larger numbers and leaving smaller numbers alone

## Slide 42: Apply Master Gain
Brightens entire image
Add constant value (in dB) to each pixel
Retrieve value ranging from -25 to 25dB from Ultrasound machine

## Slide 43: Apply Time Gain Compensation
Sound waves power decrease at square of radial distance
Results in darker regions farther away
Compensate by adding values based on distance away from transducer
Eight knobs with ranges from -20 to 20dB
  Values are set at increments of Depth / 8 (2cm)
  Linearly interpolate between each of the knobs for values for each depth

## Slide 44: Clip Dynamic Range
Monitor maps values from 0-255 linearly
At this point, image has values from 0 – 70dB
Image looks bright because most content in 50-60dB and that’s ~85% of the maximum
Clip min/max values and normalize
Optimal dynamic range found to be 30-95dB

## Slide 45: Scan Conversion
Phased array transducer creates fanning pattern
Convert from polar domain to Cartesian
Parameters:
  Initial Radius: 30px
  Final Radius: 508px
  Initial Angle: 228º
  Final Angle: 308º

## Slide 46: Noise Parameters
Gaussian filter
  Std. dev = 1.5
Perona-Malik anisotropic diffusion
  # Iter = 5, Conduction Coeff = 30, Speed of Diff = 0.2
Median filter
  Kernel size of 3
Lee filter
  Kernel size of 3
Wavelet denoising
  Two-level Daubechies-1
  Hard-thresholding, variance = 0.5
Notes: Smoothing filter applied to image

## Slide 47: Converting RF to IQ data
We get IQ data from machine, conversion already done
1. Down-mixing (shifting signal to baseband)
2. Low-pass filter (remove negative copy of signal)
3. Decimate (reduces bandwidth for storing signal)

## Slide 48: Filtering Second Harmonic
Not sure of sampling frequency, fundamental or harmonic frequency. Waiting to hear back from ultrasound rep.
IQ data contains fundamental and harmonic frequency
Down-mix the signal so that harmonic frequency is baseband
Low-pass filter to remove the fundamental frequency
This is the result without filtering harmonic frequency

## Slide 49: Tissue Harmonic Imaging (THI)
Improvement from traditional ultrasound
Transmit at one frequency (fc)
Receive at second harmonic (2 * fc)
Increases ultrasound image quality
  higher frequency = quicker changes in image, thus better res.